## BLE characteristics
Service | Characteristic | Properties | Input/Output
nstim | command | write | 5 byte array. First byte refers to command, next 4 are params.
| err | read | 1 byte.
| stim_state | read | 2 byte array. Stim state (0 off, 1 on) and stim type (0 burst, 1 dc).
dc_serv | tar_curr | read | 4 byte array. Float referring to target value of DC stim, current. Little endian format.
| base_curr | read | 4 byte array. Float referring to base value of DC stim, current. Little endian format.
| timing | read | 12 byte array. 3 sets of 4 bytes, each in little endian format. Order: Slope, Flat, Interstim
| pulses | read | 4 byte array. Little endian uint32_t. Number of pulses we are aiming to complete with DC
burst_serv | pulses | read | 4 byte array. Little endian uint32_t
| bursts | read | 4 byte array. Little endian uint32_t
| p1_curr | read | 4 byte array. Little endian float giving current of phase 1
| p2_curr | read | 4 byte array. Little endian float giving current of phase 2
| anodic | read | 1 byte array. 1 if anodic, 0 if not.
| timing | read | 20 byte array. 5 sets of 4 bytes, each set in little endian format. Order: Interstim, P1, Interphase, P2, Interburst

## Command details
Command | Code | Params | Description | Error
NOP | 0x00 | NULL | Does nothing at all | None
Start | 0x01 | NULL | Starts the stimulation. | 1 if load balancing fails. 2 if compliance check fails
Stop | 0x02 | NULL | Stops the stimulation. Should always succeed. | None
Stim type | 0x03 | Final byte in array set to 0 or 1 | Changes between burst mode and DC mode. 0 is burst, 1 is DC | 1 if stim type invalid
Anodic/Cathodic | 0x04 | Final byte in array set to 0 or 1 | 0 is anodic, 1 is cathodic | 1 if stim type invalid
Burst phase 1 | 0x05 | 4 byte array -> 32 bit unsigned | Set first phase time. Pass little endian value | None
Burst phase 2 | 0x06 | 4 byte array -> 32 bit unsigned | Set second phase. Pass little endian value in microseconds | None
Burst stim gap | 0x07 | 4 byte array -> 32 bit unsigned | Set interphase gap. Pass little endian value in microseconds. | None
Burst burst gap | 0x08 | 4 byte array -> 32 bit unsigned | Set interburst gap. Pass little endian value in microseconds. | None
Burst pulse num | 0x09 | 4 byte array -> 32 bit unsigned | Define number of pulses per burst. 0 means continuous | None
Burst burst num | 0x0A | 4 byte array -> 32 bit unsigned | Define number of bursts. 0 means continuous | None
Burst phase 1 curr | 0x0B | 4 byte array -> 32 bit float | Takes in current, sets burst phase 1 to corresponding vdac val | 1 if anodic and passing val lower than V0. Vice versa for cathodic. 2 OOB
Burst phase 2 curr | 0x0C | 4 byte array -> 32 bit float | Takes in current, sets burst phase 2 to corresponding vdac val | 1 if cathodic and passing val lower than V0. Vice versa for anodic. 2 OOB
DC ramp time | 0x0D | 4 byte array -> 32 bit unsigned | Time in microseconds for ramping in DC | None
DC hold time | 0x0E | 4 byte array -> 32 bit unsigned | Time to hold dc max current | None
DC curr target | 0x0F | 4 byte array -> 32 bit float | Takes in current, sets dc curr target to corresponding vdac val | 1 OOB